PRE-INSTALLATION REQUIREMENTS:

P1: CARESTREAM - GENERAL DATA:

2. Electrical Specification:

Three Phase Supply 100Amp (415Vac/50Hz) - 5 cores L1, L2, L3, Neutral and Earth.

Mains Impedance at Generator (item 4) <= 0.12 ohms.

Single Phase Supply 20Amp (240Vac/50Hz) - L1, Neutral and Earth.

For further electrical requirements see sheets 2-9.

3. Environmental Specification & Operating Conditions:

Temperature Range . . . . . . . . . . . . . . . . . . . . . . . 10°C to 35°C

Relative Humidity Range . . . . . . . . . . . . . . . . . . . .45% to 85% non-condensing

Equipment Heat Outputs:

Standby . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 929kcal/h = 3685 BTU/h = 1.08 kW/h

Full Power . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 4448kcal/h = 17,640 BTU/h = 5.17 kW/h

The installation of dedicated air-conditioning units in the X-Ray room is recommended.

4. Ceiling Height:

The recommended height from the finished floor to the finished ceiling is 2700mm.

The minimum ceiling height is 2350mm.

P2: ELECTRICAL REQUIREMENTS:

1. All electrical wiring interconnections etc. must be strictly in accordance with the recommendations contained within

BS7671:2001 IEE Wiring Regulations, the latest MEIGaN regulations and NHS Estates Document HTM 06-01 and

BS EN 60601-1-1:2006.

2. Electrical Components:

i) Contractor to provide and fit all the Trunking, Switchgear,Isolator & Contactor, Earth Reference Bar containment and

Emergency Pushbuttons and cables where required.

ii) Contractor to clearly label all the above items to indicate their functions.

3. Power Supply: Contractor

i) Provide Three Phase supply 100Amp (415Vac/50Hz) - 5 cores L1, L2, L3, Neutral and Earth, protected by a 63Amp

HRC SB fuse, for Generator Cabinet (item 4) - Supply exit point 'A'.

ii) Mains Impedance at item 4 <= 0.12 ohms.

iii) Provide a Single Phase supply 20 Amp (230Vac/50Hz) protected by a RCD (Type C) in the contactor box, for the

BK-200 Cabinet (item 6) - Supply exit point 'D'.

Note:- The above supplies will require 3 metre SY cable tails.

4. Start/Stop Controls must be located behind Lead Protective Screen, see sheet 3.

i) All Emergency Stop buttons must isolate the mains power to the Shimadzu FH-20 HR upon operation.

5. Electrical Contactors:

i) Provide a suitable contactors for the feed to Generator Cabinet (item 4), exiting from point 'A' (see sheet 3) leaving a

3 metre tail.

6. Earthing - Contractor to provide all Flexible Earth Cables as below:

i) Earth cable (16mm² ) from ERB to Generator Cabinet (item 4) See sheet 3 - Exit Point 'A'

ii) Earth cable (16mm² ) from ERB to Elevating Table (item 3) See sheet 3 - Exit Point 'B'

iii) Earth cable (16mm² ) from ERB to Erect Bucky (item 2) See sheet 3 - Exit Point 'C'

v) Earth cable (16mm² ) from ERB to X-Ray tube (item 1) See sheet 3 - Exit Point 'D'

P3: FLOOR REQUIREMENTS:

1. Contractor must ensure that the floor structure is capable of withstanding the weights of items 1, 2, 3 & 4 and that the

required fixing methods (see sheet 2) are achievable.

2. Floor should be suitable to withstand equipment loadings - see note P1-1.

3. Floor area under items 1, 2 & 3 must be clear from trunking, heating pipe work, electrical cabling, etc. See sheet 2 for

details.

4. Floor structure to be constructed to eliminate all vibration.

5. Check the floor under items 1, 2 & 2 is flat and level in the areas as shown at sheet 2.

6. Existing floor needs to be structurally validated (if required).

7. Floor trunking - Contractor to provide Floor Trunking as below:

i) Floor trunking and excavation requirements see sheet 2.

ii) Final section of vinyl flooring to be installed & welded after equipment installation.

iii) In the event that the vinyl flooring is laid over the trunking two pull cables need to be fitted inside the trunking before

hand. Trunking to be debarred and clear of debris.

iv) Provide floor protection/covering ready for the delivery and installation of the X-Ray equipment.

P4: CEILING REQUIREMENTS:

THE RECOMMENDED HEIGHT FROM THE FINISHED FLOOR TO THE FINISHED CEILING IS 2700mm.

THE MINIMUM CEILING HEIGHT IS 2350mm.

1. Ceiling Cable Tray - Contractor to provide Ceiling Cable Tray as below (if required):

i) 100mm wide cable tray above suspended ceiling at the lowest possible level, e.g upper side of unistrut.

ii) All cable tray runs to be kept to a minimum as indicated, e.g point to point.

iii) If the indicated route cannot be achieved then please advice MED XRAY of the alternative route to enable:

a) Correct X-Ray equipment configuration, e.g system cable lengths.

b) Update room drawings.

c) X-Ray equipment installation timescale.

P5: RADIATION PROTECTION:

The radiation protection requirements for the X-Ray room are the responsibility of the Trust RPA.

All Released Version Drawings are required to be signed off by the Trust RPA.

1. Room Warning Lights:

i) Room Warning Lights and cabling to be supplied and fitted by Trust/Contractor.

ii) Wording for the Room Warning Lights to be supplied and approved by the Trust RPA.

iii) Final position of Room Warning Lights to be supplied and approved by the Trust RPA.

2. Room Warning Lights - lamp type:

i) Upper section (yellow) - lamp type fluorescent or filament.

ii) Lower section (red) - lamp type filament.

3. All SY cables for Room Warning Lights to be provided as follows leaving a 3 meter tails:

i) Warning Lights to Generator Cabinet (item 4) - Exit point 'A'.

4. Lead Protective Screen:

i) Final Layout/Screen design to be agreed and signed off by the Trust before manufacture.

P6. GENERAL DRAWING NOTES:

1. All drawings have been produced based on the current room information.

2. Any furniture, fixings etc. slready installed are shown on the drawings for illustration purposes only.

3. Please consult MED XRAY of any future room changes to ensure that the changes have no effect on the correct

operation of the X-Ray equipment.

4. All Released Version Drawings are to be reviewed and signed off by the following:- Trust (users), Trust RPA, Contractor.